## Lessons Learned
ID | Area | Symptom / What Happened | Root Cause | Fix / Resolution | Lesson Learned
1 | VISA | Assumed TekVISA was the Python driver to program against. | TekVISA is a VISA backend (a DLL), not a Python API. The Python driver is a separate package, tm_devices. | Use tm_devices as the Python driver; TekVISA (or NI-VISA) sits underneath it via PyVISA. | Know the layers: script -> tm_devices -> PyVISA -> VISA backend -> instrument.
2 | VISA | pyvisa raised VI_ERROR_INV_OBJECT (-1073807346) from visa32.dll. | NI-VISA and TekVISA were both installed. Both register a visa32.dll and the IVI conflict manager ended up pointing at no usable provider. | Uninstall one of the two VISA implementations, then reboot. | Never install two VISA implementations on the same machine.
3 | VISA | NI-VISA was listed in Add/Remove Programs but nothing worked. | Only headers, passports and documentation had installed. The actual runtime (nivisa64.dll) was never present. | Reinstall NI-VISA choosing the full runtime (plus NI-MAX), not just the SDK/headers. | 'Listed as installed' does not mean working. Verify the runtime DLL actually exists on disk.
4 | VISA | Removing NI-VISA entry-by-entry from Add/Remove Programs failed. | NI-VISA installs ~20 interdependent components. | Use NI Package Manager (nipkg.exe). It must be run as Administrator. Preview first with 'remove --simulate'. | Use the vendor's own package manager to remove vendor software; preview destructive operations before running them.
5 | VISA | TekVISA's Instrument Manager saw the scope and Talker/Listener answered *IDN?, but Python could not open a VISA session at all. | TekVISA 4.1.1 installs a WORKING 32-bit VISA (SysWOW64\visa32.dll = Tektronix) but only an IVI stub for 64-bit (System32\visa32.dll = IVI Foundation). Tektronix's own tools are 32-bit apps; Python was 64-bit. | Install TekVISA 4.2.0 or newer, or install 64-bit NI-VISA. | If the vendor's GUI works but your code does not, suspect a bitness mismatch. Check the CompanyName on the VISA DLLs in System32 vs SysWOW64.
6 | VISA | Considered downgrading tm_devices so it would accept the older TekVISA 4.1.1. | There is NO TekVISA version check anywhere in the tm_devices code. The '>= 4.2.0' is only a documentation recommendation. | Do not downgrade tm_devices. Fix the VISA install instead. | Confirm a stated 'requirement' is actually enforced before building a workaround for it.
7 | USB | Windows never detected the scope on USB, across several attempts and reboots. | Wrong cable/port. A USB-C cable was used. The scope's PC-control port is USB-B (the 'USB Device' port); the others are host ports for flash drives. | Use a USB-B to USB-A cable into the scope's USB Device port. | Only the device port presents the scope to a PC. Also beware charge-only cables, which carry power but no data.
8 | USB | Could not tell whether the USB cable was carrying data at all. | No way to distinguish a dead cable from a driver problem by guessing. | Snapshot the Windows USB device list, unplug, replug, and diff it. If NOTHING appears or disappears, the cable/port carries no data. | Prove the physical layer with OS tools before debugging any software.
9 | USB | Needed a reliable way to confirm the scope was enumerated. | n/a | Get-PnpDevice -PresentOnly, filtered on InstanceId matching VID_0699 (Tektronix's USB vendor ID). Status 'OK' means enumerated and driver bound. | VISA cannot find a device that the operating system has not enumerated.
10 | USB | pyvisa-py could not talk to the scope over USB. | USBTMC requires a real VISA backend on Windows; pyvisa-py is not sufficient. | Install NI-VISA or TekVISA for USB work. | USB requires a full VISA. Only LAN can fall back to pyvisa-py.
11 | LAN | The scope reported IP 169.254.8.134 and the PC could not reach it. | 169.254.x.x is a link-local (APIPA) address, self-assigned when there is no DHCP server. It means the scope is on a direct cable, not a router. | Put the PC's Ethernet adapter on the same link-local subnet, e.g. New-NetIPAddress -InterfaceAlias Ethernet -IPAddress 169.254.8.100 -PrefixLength 16 (as Administrator). | A 169.254.x.x address means no DHCP. Both ends must be on the link-local subnet to talk.
12 | LAN | Ping to the scope succeeded, but the script still could not talk to it. | Ping only proves the network layer. It says nothing about whether the instrument is listening for SCPI. | Test the SCPI transport separately (port scan, then an actual *IDN? query). | Ping success does not mean the instrument is reachable for SCPI.
13 | LAN | Did not know which LAN protocols the scope was actually offering. | n/a | Port-scan it: Test-NetConnection <ip> -Port 111 (VXI-11), -Port 4880 (HiSLIP), -Port 4000 (raw Socket Server). Only 'True' ports are available. | Port-scan the instrument to find out what it really offers instead of assuming.
14 | LAN | tm_devices with pyvisa-py failed over VXI-11 (VI_ERROR_RSRC_NFOUND, instantly) and over HiSLIP (VI_ERROR_RSRC_NFOUND after a 5 second timeout). | pyvisa-py's VXI-11 and HiSLIP implementations could not drive this scope. | Use the scope's raw Socket Server (bench_socket.py), or install a real VISA. | pyvisa-py is not a drop-in replacement for a real VISA on LAN. Do not assume it will work.
15 | LAN | Connections to the Socket Server port were refused, even though the port and settings looked right. | On the scope, the Socket Server 'Protocol' was set to 'None', which turns the server OFF and leaves the port closed. | On the scope: Utility -> I/O -> Socket Server -> set Protocol = Terminal. Confirm with Test-NetConnection that the port now opens. | 'None' does not mean 'raw mode'; on this scope it means the server is off. Verify the port actually opens after changing a setting.
16 | Simulator | pyvisa-sim could not return a binary waveform. | A binary CURVe? block contains arbitrary bytes including 0x0A, which collides with the newline read-terminator of the text-based simulator. | Use ASCII encoding (DATa:ENCdg ASCii) so the curve comes back as comma-separated text. | Binary transfers cannot pass through a newline-terminated text simulator.
17 | Simulator | tm_devices could not connect to pyvisa-sim at all. | On connect, tm_devices calls read_stb() (read status byte) and clear() (device clear) and drains the output buffer. pyvisa-sim implements none of these and raises NotImplementedError. | Keep the offline/simulated path on plain PyVISA; use tm_devices only against real hardware. | Split the code: scripts/sim (offline, pyvisa-sim) vs scripts/real (hardware, tm_devices).
18 | Socket / Parsing | Replies from the scope contained the echoed command and a prompt character. | The Socket Server's Terminal mode is interactive: it echoes what you send and adds a prompt. | Clean each reply: drop the echoed command, strip prompt characters and whitespace. | Terminal mode is meant for humans. Replies must be cleaned before they can be parsed.
19 | Socket / Parsing | Requesting 10,000 waveform points returned exactly the same data as requesting 1,000. More points changed nothing. | The reply parser kept only the LAST line of the response (correct for a single-value query). A long waveform is wrapped across many lines, so almost all the data was being discarded. | For the curve, read the RAW reply, strip ALL whitespace (which also heals a number split across a wrap boundary, e.g. '18\n0' -> '180'), then split on commas. | A parser that works on short replies can silently truncate long ones. Always verify that asking for more data actually returns more data.
20 | Waveform | The captured time span did not match the time window set on the scope's display. | DATa:STARt / DATa:STOP select samples by INDEX from the whole record, not by time. The record is usually far longer than the displayed window, so the first N samples come from the start of the record. | Convert the desired times into sample indices using the preamble (XZERO, XINCR), or transfer the entire record. | Waveform transfer windows are index-based, not time-based.
21 | Waveform | The raw curve values were meaningless integers. | The scope returns integer ADC codes, not volts. | Scale with the WFMOutpre preamble: volts = (code - YOFF) * YMULT + YZERO ; seconds = XZERO + (n - PT_OFF) * XINCR. | Always scale via the preamble. Confirm PT_OFF handling against the real instrument.
22 | Plotting | The plotted waveform's amplitude disagreed with the Vpp reported in the summary. | The plot picked one sample per column, which aliased the signal and missed the actual peaks. | For each column, draw the min..max of every sample that falls in it (an envelope). | Downsampling for display must preserve peaks, or the picture lies about the data.
23 | Verification | Read-back comparison reported failures on values that were actually correct. | The instrument formats values differently from what you sent: 500.0000E-3 for 0.5, RIS for RISE. | Compare numbers numerically with a tolerance; compare strings case-insensitively and allow the instrument's abbreviated form. | Never compare instrument replies with a plain equality check.
24 | Verification | Some tm_devices command nodes had no .query() method (e.g. CH:SCAle). | Those nodes are modeled as write-only in the generated API. | Fall back to a raw query built from the node's cmd_syntax attribute. | Do not assume a generated API is symmetric; check that a read path exists before relying on it.
25 | Windows / Python | The script crashed with UnicodeEncodeError when printing. | The Windows console uses cp1252. Printing Unicode (plus/minus signs, check marks, box-drawing characters) fails. | Keep ALL printed output plain ASCII. | The console's encoding limits what you can print. Do not put Unicode in program output on Windows.
26 | Windows / Python | 'set VAR=value' did not set the environment variable. | That is cmd.exe syntax, not PowerShell. | In PowerShell use: $env:SCOPE_HOST = "169.254.8.134" | Know which shell you are in; the syntax differs.
27 | Windows / Python | Activating the virtual environment was blocked by the execution policy. | PowerShell blocks running Activate.ps1 by default. | Skip activation and call the venv's interpreter directly: .\.venv\Scripts\python.exe script.py | You never actually need to 'activate' a venv; calling its python.exe directly always works.
28 | Process | Repeatedly debugged the wrong layer of the stack. | No systematic order to the investigation. | Work bottom-up: cable -> OS device detection -> VISA -> PyVISA -> tm_devices -> your script. Do not debug a layer until the one below it is proven. | Isolate layers and prove them in order. Most time was lost debugging software while the cable was dead.
29 | Process | 'It failed' was not enough information to act on. | The error text was being skimmed rather than read. | Read the error CODE and the time-to-fail. VI_ERROR_INV_OBJECT = the VISA install itself is broken. VI_ERROR_RSRC_NFOUND = VISA works but that protocol/address is not answering. An instant failure is a refusal; a 5-second failure is a timeout. | The error code and the speed of failure both carry diagnostic information.
30 | Process | All progress was blocked while the VISA installation was broken. | Every path depended on a working VISA and a working tm_devices. | Wrote bench_socket.py: talks to the scope's raw Socket Server using only the Python standard library. No VISA, no tm_devices, no pip installs. | Keep a fallback path that depends on as little as possible. It is what kept working when everything else failed.