# Oscilloscope Capture - Test Developer Guide

This guide is for the TEST DEVELOPER building a TestStand sequence. It explains the flow of logic - which functions to call, in what order, and why - to capture a waveform off the oscilloscope and check it, and it lists each function with its parameters and return value.

The functions live in teststand_api.py (the oscilloscope one). Each one takes and returns simple values (string, number, boolean), so you bind them straight to sequence variables.

# Index

Index	1

1. The big picture	2

2. The flow of logic	2

2.1 Step by step	2

2.2 The two rules	3

2.3 About the trigger position (5%)	3

3. Function reference	3

3.1 Connection	3

3.2 Setup	4

3.3 Capture	5

3.4 Measurements (put LIMITS on these)	5

3.5 Saving	6

4. Applying test limits	7

5. Saving results (the path rules)	7

6. Pass/fail, errors, and single=False	7

7. Common problems	8

8. Quick reference - the whole sequence	8

# 1. The big picture

The most important thing to understand: the WAVEFORM IS ALREADY THERE before your capture sequence runs. It is produced earlier in the test - by the function generator, or by the unit under test - and it is sitting on the oscilloscope by the time you read it.

Because the waveform already exists, the scope does NOT need to arm and wait for a trigger. Your sequence simply connects to the scope, applies the setup, and READS the record that is already on screen. That is the whole idea.

So the logic is: generate the waveform first (a separate, earlier step), then run the capture sequence below to read and check it.

# 2. The flow of logic

These are the steps of the capture sequence, in order. The waveform-generation step happens BEFORE step 1.

--- the waveform is generated earlier in the test ---

1. connect("169.254.8.134")        -> String   (scope identity)

2. configure("bench_full", "1,2")  -> Boolean  (pass/fail)

3. capture("1,2", single=False)    -> Boolean  (read the current record)

5. save_png("C:\\results\\run.png") -> String   (path written)

6. disconnect()                    -> Boolean

## 2.1 Step by step

Step 1 - connect(host)

Opens the connection to the scope and returns its identity string. This is the first step; nothing else works before it. If the scope cannot be reached, it raises an error (TestStand shows a step error).

Step 2 - configure(setup_name, channels)

Applies a named setup to the scope (vertical, horizontal, trigger), then reads every setting back and checks it landed. It returns True only if every setting verified - bind that to a pass/fail step. The setup name is a file in the configs folder without the '.json' (for example "bench_full"). Pass "1,2" for the channels you want, or "" to use the channels the setup already defines.

Step 3 - capture(channels, single=False)

Reads the waveform that is already on the scope into memory, so the measurement and save steps can use it. single=False is the key: it means "read the record that is already there". It returns True if data came back for every channel you asked for. Do NOT use single=True here (see section 6).

Step 4 - save (save_png, save_csv)

Writes the captured waveform to a file and returns the path(s) written. Bind the return to a String and attach it to the report. See section 5 for how to write the path.

Step 5 - disconnect()

Closes the connection. Put this last, in the Cleanup section, so it runs even if a step above fails.

## 2.2 The two rules

Step 3 uses single=False. The waveform is already on the scope, so single=False just reads it. single=True would arm and wait for a NEW trigger that never comes, and time out.

configure() runs once, in step 2. Do not reconfigure the scope later in the flow.

## 2.3 About the trigger position (5%)

The scope setup uses a 5% horizontal trigger position. That means 5% of the record sits BEFORE the trigger point and 95% sits after it. You do not set this in your sequence - it is part of the named setup - but it is worth knowing: almost the whole record is the post-trigger window, so a signal that appears at or just after the trigger is captured with room to spare.

# 3. Function reference

Every function you call in the flow, with its parameters and return value. All values are primitives, so they bind straight to TestStand variables.

## 3.1 Connection

connect(host: str, port: int = 4000) -> str

Opens the connection to the scope and returns its identity string. Call this first.

Parameters:

host (str) - the scope IP address, e.g. "169.254.8.134"

port (int) - Socket Server port. Default 4000.

Returns: str - the identity string.

Errors: Raises RuntimeError if the scope cannot be reached (off, wrong IP/port, wrong network).

disconnect() -> bool

Closes the connection. Put it last, in Cleanup.

Returns: bool - always True.

identify() -> str

Returns the scope's identity string on an already-open connection.

Returns: str - the identity string.

Errors: Raises RuntimeError if not connected.

## 3.2 Setup

list_setups() -> list[str]

The names of the setups available in the configs folder.

Returns: list of str, e.g. ['bench', 'bench_full', 'default'].

configure(setup_name: str = "bench_full", channels: str = "") -> bool

Applies a named setup to the scope, then reads every setting back and checks it landed. Run this once, in Setup.

Parameters:

setup_name (str) - the name of a file in the configs folder without the .json, e.g. "bench_full".

channels (str) - "1,2" for specific channels, or "" to use the channels the setup defines.

Returns: bool - True only if every setting read back correctly.

Errors: Raises ValueError if the setup name is unknown; RuntimeError if not connected.

Good to know: Call get_config_report() afterwards for the PASS/FAIL detail.

get_config_report() -> str

The PASS/FAIL table from the last configure(), as text. Bind to a String and attach to the report.

Returns: str - a multi-line table, one line per setting.

## 3.3 Capture

capture(channels: str = "1", points: int = 0, single: bool = False, timeout_s: float = 120.0) -> bool

Reads one or more channels off the scope into memory for the measurement and save steps.

Parameters:

channels (str) - "1" for one channel, "1,2" for several.

points (int) - samples to transfer. 0 (default) = the whole record the setup defines.

single (bool) - Use False - read the record already on the scope. True waits for a NEW trigger (not what you want here).

timeout_s (float) - only used when single=True.

Returns: bool - True if data came back for every requested channel.

Errors: Raises RuntimeError if not connected. (single=True can raise TimeoutError.)

captured_channels() -> list[int]

Which channels actually returned data on the last capture().

Returns: list of int, e.g. [1, 2].

## 3.4 Saving

save_csv(path: str) -> str

Saves the captured data as CSV.

Parameters:

path (str) - absolute path with the file name; see section 5.

Returns: str - the path(s) written, semicolon separated.

Errors: Raises RuntimeError if nothing has been captured yet.

save_png(path: str) -> str

Saves a PNG plot of the captured data.

Parameters:

path (str) - absolute path with the file name; see section 5.

Returns: str - the path(s) written, semicolon separated.

Errors: Raises RuntimeError if nothing has been captured yet.

# 4. Applying test limits

The measurement functions in section 3.4 are where the test actually passes or fails on the signal. Bind one to a Number step and set the limits on that step.

Example: to check a 2 V peak is within 5 percent, put get_vmax(1) on a Number step with limits 1.9 to 2.1.

If you ask for a channel that was not part of the last capture(), the measurement raises an error - capture that channel first, or check captured_channels().

# 5. Saving results (the path rules)

save_png(path) and save_csv(path) return the path(s) they wrote. A few rules for the path:

Use DOUBLE backslashes in Windows paths: "C:\\results\\run.png". A single backslash can be read as an escape code and corrupt the path.

Include the file NAME and extension, not just a folder: "...\\run.png", "...\\run.csv".

Use an ABSOLUTE path starting with a drive letter. A relative path lands in TestStand's working folder and seems to vanish.

With one channel you get exactly the file you named. With several channels you get one file per channel plus a combined file - for example run_CH1.png, run_CH2.png and run_joint.png. The return value tells you the exact paths (semicolon separated).

# 6. Pass/fail, errors, and single=False

How TestStand should read the results:

A function that returns False = the TEST FAILED (a setting did not verify, or a channel returned no data). Bind these boolean returns to a pass/fail step.

A function that RAISES = something is BROKEN (scope unreachable, wrong channel). TestStand shows a step error with the message.

The single most common mistake is using single=True in step 3. Here is why it is wrong:

single=False means "read the record that is already on the scope". That is what you want, because the waveform was generated before your sequence.

single=True means "arm the scope and WAIT for a new trigger, then read". Since the waveform already happened, no new trigger comes, and the step times out with an error.

Unless you are deliberately catching a brand-new event that has not happened yet, always use single=False.

# 7. Common problems

configure() returns False

One or more settings did not read back. Call get_config_report() to see which line FAILED - often the scope rounded a value to the nearest legal one, which the readback will show.

capture() returns False, or only one file is saved

A channel returned no data - it was not displayed or not acquiring on the scope. Make sure the signal is present and the channel is on. captured_channels() shows what came back.

capture(single=True) times out

You used single=True but the waveform is already on the scope. Switch to single=False (section 6).

A measurement raises "No captured data for CHn"

That channel was not in the last capture(). Capture it first.

The saved file cannot be found

You used a relative path. Use an absolute path with double backslashes, ending in the file name (section 5).

connect() raises "Cannot reach the scope"

The scope is off, on a different network, or the Socket Server is not on. Confirm the IP address and that the scope is reachable.

# 8. Quick reference - the whole sequence

Setup:

connect("169.254.8.134")            -> String

configure("bench_full", "1,2")      -> Boolean   (fail the step if False)

get_config_report()                 -> String    (optional, attach to report)

Main:

capture("1,2", single=False)        -> Boolean   (fail the step if False)

save_png("C:\\results\\run.png")     -> String

save_csv("C:\\results\\run.csv")     -> String

Cleanup:

disconnect()                        -> Boolean

Remember: the waveform is generated BEFORE this sequence. This sequence only reads and checks it.